## Slide 1: Energy calibration at Dalton Cumbria Facility
Energy calibration of a 2.5MV Pelletron at dcf

## Slide 2
Energy calibration of the 2.5 MV pelletron
Yield measurement of resonance strengths in 27Al(p,γ)28Si due to the well-known strength values
Proton beam onto a 27Al target and a NaI γ -ray detector

## Slide 3: Resonances in 27Al(p, γ)28Si
Ep | Γ
632.23 ± 0.04 keV | 6.7 ± 0.5 eV
991.86 ± 0.017 keV | 70 ± 10 eV
1213.08 ± 0.06 keV | < 100 eV
1587.49 ± 0.08 keV | < 200 eV
1799.75 ± 0.09 keV | 450 ± 60 eV

## Slide 4: Experimental set up
Aluminium target
Water cooling
Actuator
NaI detector

## Slide 5: Experimental method
Instead of changing the proton beam parameters a bias voltage was applied to scan through the resonance range.
Initiate the beam with nominal energy
Set up a bias between the target and collimator by 300 V for electron suppression
Beam on target to collect a spectrum with a NaI detector
Save and reset spectrum measurement
Change the target bias voltage to continue the resonance scan

## Slide 6: Data fitting
From the 27Al(p, γ)28Si reaction we want to measure gamma emission of about 1.78 MeV so an appropriate region of interest has been selected.
The spectra are saved with the use of Ortec MAESTRO in an ASCII format to allow analysis on raw data.
Chanel
Count
1.78 MeV

## Slide 7: Detector calibration
The NaI detector was calibrated by correlating common background radiation peaks with the gamma energy.
Calibration has been performed at DCF before the first experimental run.

## Slide 8: Yield calculation
Due to unstable current output the current values have been integrated.
A Gaussian and a first-degree polynomial is used to find the area under the emission peak.

## Slide 9: 27Al(p, γ) at 991 keV
Fixed
Fixed
Fixed
Free
eV
± 30 eV
eV
Uncertainty scaled by 2.7 Χ2/N = 1.1
 |  | Χ2/N
269 ± 29 eV | 60 eV | 1.105
258 ± 30 eV | 70 eV | 1.111
246 ± 31 eV | 80 eV | 1.124
 |  | Χ2/N
269 ± 29 eV | 60 eV | 1.105
258 ± 30 eV | 70 eV | 1.111
246 ± 31 eV | 80 eV | 1.124

## Slide 10: Second measurement at 991 keV
Fixed
Fixed
Fixed
Free
eV
eV
eV
 |  | Χ2/N
196 ± 32 eV | 60 eV | 1.241
191 ± 34 eV | 70 eV | 1.294
178 ± 36 eV | 80 eV | 1.352
 |  | Χ2/N
196 ± 32 eV | 60 eV | 1.241
191 ± 34 eV | 70 eV | 1.294
178 ± 36 eV | 80 eV | 1.352
Uncertainty scaled by 2.7 Χ2/N = 1.2

## Slide 11: 27Al(p, γ) at 1213 keV
Fixed
Fixed
Fixed
Free
eV
27 eV
eV
 |  | Χ2/N
183 ± 27 eV | 100 eV | 1.051
263 ± 19 eV | 10 eV | 0.985
191 ± 34 eV | 84 ± 22 eV | 1.0378
153 eV | 116 ± 23 eV | 1.090
221 eV | 56 ± 22 eV | 1.002
 |  | Χ2/N
183 ± 27 eV | 100 eV | 1.051
263 ± 19 eV | 10 eV | 0.985
191 ± 34 eV | 84 ± 22 eV | 1.0378
153 eV | 116 ± 23 eV | 1.090
221 eV | 56 ± 22 eV | 1.002
Uncertainty scaled by 1.4 Χ2/N = 1.0
from 992 keV measurement

## Slide 12: 27Al(p, γ) at 1587 keV
Fixed
Fixed
Fixed
Free
eV
51 eV
48 eV
eV
Literature value
 |  | Χ2/N
172 ± 51 eV | 32 ± 48 eV | 0.647
194 ± 14 eV | 10 eV | 0.617
1 ± 550 eV | 200 eV | 1.696
191 ± 34 eV | < 20 eV | 0.688
153 eV | 26 ± 15 eV | 0.679
221 eV | 3 ± 8 eV | 0.885
 |  | Χ2/N
172 ± 51 eV | 32 ± 48 eV | 0.647
194 ± 14 eV | 10 eV | 0.617
1 ± 550 eV | 200 eV | 1.696
191 ± 34 eV | < 20 eV | 0.688
153 eV | 26 ± 15 eV | 0.679
221 eV | 3 ± 8 eV | 0.885
Χ2/N = 1.0
from 992 keV measurement

## Slide 13: 27Al(p, γ) at 1799 keV
Fixed
Fixed
Fixed
Free
eV
64 eV
eV
Χ2/N = 1.0
 |  | Χ2/N
362 ± 29 eV | 10 eV | 1.127
172 ± 64 eV | 200 eV | 1.548
191 ± 34 eV | 157 ± 34 eV | 1.478
153 eV | 180 ± 33 eV | 1.534
221 eV | 137 ± 33 eV | 1.4
 |  | Χ2/N
362 ± 29 eV | 10 eV | 1.127
172 ± 64 eV | 200 eV | 1.548
191 ± 34 eV | 157 ± 34 eV | 1.478
153 eV | 180 ± 33 eV | 1.534
221 eV | 137 ± 33 eV | 1.4
from 992 keV measurement

## Slide 14: 27Al(p, γ) at 632 keV
Fixed
Fixed
Fixed
Free
eV
34 eV
39 eV
eV
Literature value
Uncertainty scaled by 1.4 Χ2/N = 1.2
 |  | Χ2/N
457 ± 266 eV | 95 ± 288 eV | 1.153
466 ± 2 eV | 6.7 eV | 1.245
191 ± 34 eV | 376 ± 40 eV | 1.029
153 eV | 397 ± 39 eV | 1.027
221 eV | 356 ± 40 eV | 1.034
 |  | Χ2/N
457 ± 266 eV | 95 ± 288 eV | 1.153
466 ± 2 eV | 6.7 eV | 1.245
191 ± 34 eV | 376 ± 40 eV | 1.029
153 eV | 397 ± 39 eV | 1.027
221 eV | 356 ± 40 eV | 1.034
from 992 keV measurement

## Slide 15: Reproducibility
Two 991 keV measurements have been performed to check for reproducibility.
Both the data and fits are close to each other which suggests good reproducibility

## Slide 16: Stability
Voltage drift of 480 V during the measurement

## Slide 17: Second stability run
Voltage drift of 400 V during the measurement

## Slide 18: Calibration fit
TV: Terminal Voltage
EV: Extraction Voltage
keV/kV
keV
Residual (r)